Patchit

The Manual

Introduction:

This manual describes the operation of the Patchit data

acquisition program.  The current revision of this program is 2.30.

This software and manual are copyright 1993 through 1989 by George

B. Grant, Somerville, MA  02144.

Compatibilities:

Patchit runs on IBM PC compatible computer systems.  It

operates entirely in graphics mode.  As such, it requires a

compatible graphics display adapter.  Patchit supports all major

display adapters (Hercules, CGA, EGA, VGA) and has drivers for

several other graphics display adapters including Super VGA.

Patchit works with a Microsoft compatible mouse.

Patchit is a data acquisition program.  It therefore requires

a data acquisition card.  Patchit should work with all cards in the

Data Translation DT2821 family (except the DT2829) and in

particular, has been thoroughly tested with both the DT2828 and

DT2821 data acquisition cards.  Patchit also works with both the

Tecmar Labmaster and Labmaster DMA data acquisition cards.   For

configuring these boards, see the appendix at the end of this

manual.  As of this time, these are the only data acquisition

adapters supported by Patchit.

Starting Up:

Patchit runs entirely in graphics (not text) mode.  In order

for it to run, a graphics driver for the graphics adapter that is

installed in the computer must be available in the directory in

which the Patchit executable file is.  These drivers have the file

extension ".bgi".  The drivers required for the following screen

types are as follows:

Display Adapter                         Display Driver

Hercules (Hercules Graphics Adapter)         herc.bgi

CGA (Color Graphics Adapter)            cga.bgi

EGA (Enhanced Graphics Adapter)              egavga.bgi

VGA (Video Graphics Adapter)            egavga.bgi

Super VGA                          svga16m.bgi

(Supports several common SVGA cards and the VESA standard

SVGA interface)

IBM 8514 Graphics Adapter                    ibm8514.bgi

AT&T 400-Line Graphics Adapter               att400.bgi

3270 PC Graphics Adapter                pc3270

By default, Patchit will automatically try to detect the

proper graphics driver for the computer.  If the proper driver is

not found by Patchit (in the same directory as Patchit) when it is

started it will end with an error message.  Otherwise the program

will start and Patchit's main screen will appear.

The Screen:

Patchit's main screen is divided into three areas.  The menu

bar is on the top of the screen, it allows the user to access all

of Patchit's commands.  The right-most portion of the menu bar

shows the amount of free conventional (real mode) memory available

for Patchit's use.  The lower part of the screen shows the status

line.  This line displays information about the current status of

the program (the amount of data sampled, the current parameter file

and the last data file saved).  The data display area fills the

screen between the menu bar and the status line.

Help:

On-line help is available throughout Patchit.  Just type the

F1 key and a relevant help message will pop up.  The help file

"Patchit.hlp" must be in the same directory as the Patchit

executable file.

Accessing the Menus:

The menu bar has a series of pull down menus.  A menu (e.g.

"File") can be accessed in a number of ways.  1) A menu can be

pulled down by clicking on the desired menu entry with the mouse.

2)  A menu can be pulled down by holding down the ALT key while the

selection letter of the desired menu is typed (the selection letter

is usually the first letter of the word and is shown underlined and

in a different color from the rest of that word (unless the current

graphics display only supports two colors, or the color

configuration has been messed with)).  For example, to pull down

the File menu, type ALT-F.  3)  A menu can be pulled down by first

activating the menu bar with the F10 or "/" key.  The file menu

will drop down, but not be selected.  The selection letter of any

of the main menu items can then be typed in order to access that

particular menu.

Certain functions on the individual menus can be accessed

directly, without pulling down the menus, by typing a hot-key (the

hot-keys are shown to the left of a menu entry, on its menu).  For

instance, the F7 key will cause Patchit to read a parameter file if

pressed when the menu bar is not activated.

The Menus:

File Menu:

Read Parameter File:  Selecting this option allows the user to load

a new set of parameters into Patchit.  The parameter sets determine

how stimulus will be generated and how data will be recorded.

Parameters are stored in files called "parameter files".  If there

are parameter files in the current directory, a selection box will

appear listing the available parameter files.  A parameter file may

be cursor selected from this box (with the mouse or keyboard) or

else the name of a parameter file may be typed in (do not type the

".prm" extension).  To change directories, type the F2 key, or

click the left mouse button in the upper, directory, window and

type the pathname of the desired directory.  The list of parameter

files will be updated for this new directory.

Write Parameter File:  Selecting this option allows the user to

store the current parameter set to a parameter file. The name of

the parameter file should be typed without the extension ".prm"

(this is automatically appended by the program).

Save Data File:  Selecting this options allows the user to save the

latest set of data that has been recorded.  When the data is saved,

a complete copy of the parameters used to record this data is also

stored along with it.  You must use this option to save the data

that you have just recorded.  If you do not, then the data recorded

by Patchit will be overwritten the next time you acquire data with

Patchit.  The only exception to this is if you have configured

Patchit to record data directly to your disk (see below,

File:Options:Record to Disk).

Options:  Selecting this option allows the user to control several

very important aspects of Patchit's operation.  These are:

Record to Disk:  Selecting this toggles the state of this feature.

When this is enabled, Patchit will store each record as it is

acquired to a disk file.  When you begin sampling data, Patchit

will prompt you for the name of the file into which the data will

be saved.  Prior to writing each record to this disk file, Patchit

will check that there is sufficient disk space available.  You will

be warned if there is not.

Comments in Files:  Selecting this toggles the state of this

feature.  When this is enabled, you will be prompted for a comment

when you save data to a disk file.  The comment can be of any

length and is useful for recording pertinent facts about the data.

Okay:  Selecting this causes any changes that you have made to the

state of these options to be saved.

Cancel:  Selecting this causes any changes that you have made to

the state of these options to be discarded.

DOS Shell:  When this is selected the user temporarily exits to DOS

to perform simple commands.  In order to allow this, Patchit

"spawns" a copy of the DOS command interpreter.  Note that Patchit

is still loaded in memory so there will be less DOS memory

available to execute a command.  Type "EXIT" at the DOS command

line ("C:\>") to return to Patchit.

About:  When this is selected another the user can find out about

Patchit.

Patchit:  Selecting this causes information about Patchit to

appear.  This information is the current revision level, a

copyright notice and information about its author (me!).

Memory:  Selecting this causes information about the memory on your

computer and Patchit's use of it to appear.  There are three types

of memory used by Patchit.

Conventional Memory:  This is the memory used by all DOS real mode

programs, it is limited to a maximum of 640 Kilobytes (KB).

Patchit will show you the total amount of conventional memory on

your computer and the amount that is currently available.

XMS Memory:  This is memory above the 1 Megabyte boundary that can

be accessed by Patchit if an XMS driver (such as Microsoft's

HIMEM.SYS) is loaded.  XMS memory allows Patchit to store a great

deal of data without filling up the DOS (conventional) memory area.

If XMS is available Patchit will store acquired data in it unless

you have disabled this feature (see below, Config:Other Options:XMS

Data Storage).

Disk Memory: This is the memory available on the disk.  Patchit

shows the total disk space and the available amount on the current

disk.

Okay:  Selecting this closes this menu.

Quit Program:  Selecting this allows the user to exit Patchit.

Display Menu:

Clear Screen:  Selecting this causes all data display windows to be

cleared.

Refresh:  Selecting this causes any data that has been recorded by

Patchit to be redisplayed in the data display windows.  Data

recorded on channels that have not been enabled for display will

not be displayed.  If you have recorded data directly to the disk

(see above, File:Options:Record to Disk) then Patchit will not be

able to refresh the screen.

Parameters:  Selecting this causes Patchit to display some of the

currently selected parameters.  This does not show all of the

parameters in effect and is in need of work.

Windows:  When this option is selected the user is prompted for the

number of display windows that will subdivide the data display

area.  Up to four such windows may be used.  If more than one is

selected the screen will clear and the user can then specify the

layout of these windows.  Lines in the data display area show the

limits of the windows and can be moved about with the cursor keys

or mouse.  Typing the space bar causes the layout/orientation of

the windows to be changed.  When you are satisfied with the layout,

hit the Enter key.  Windows are number in ascending order from left

to right and top to bottom.

Note:  If the user decreases the number of windows, then any data

that was configured to have been displayed in a window that no

longer exists (greater than the number of windows) will be

displayed in window 1.

Sample Menu:

Go:  Selecting this causes Patchit to begin acquiring data.

Stimuli will be generated on any enabled output channels and data

will be acquired on any enabled input channels.  Simultaneously.

This process may be interrupted by hitting a key.

Continue:  Selecting this causes Patchit to begin acquiring data

from the record at which it had been previously interrupted.  The

user has the option of continuing from a record earlier than that

at which data acquisition was previously interrupted.  That is, if

you interrupted the data acquisition cycle after record #8 because

of a big glitch in the recording (due to a power line spike of

course) you could continue recording at record #6, overwriting

records 6 - 8 with new data.

Number of Records:  Selecting this allows the user to specify the

number of records that Patchit will record.  Patchit records data

in blocks that are called "records".  A record can contain multiple

channels of data.  The number of records is limited to ~64000, but

is effectively limited by memory.  If you select a value and

Patchit finds that there is insufficient memory on your computer,

then Patchit will complain and reset the value.  If you need to

record such large amounts of data, you should enable

File:Options:Record To Disk (and have a big disk).

Length of Record:  Selecting this option allows the user to specify

the length of a record (in milliseconds).  Each record is limited

in size to a maximum of ~32000 points.  The number of points is

equal to ((the number of input channels) * (the length of the

record)) / (the sampling rate).  The number of points in the record

is shown under Display:Parameters.  If you choose a value for the

length of the record that exceeds the maximum number of points

allowable, or the amount of memory available, Patchit will complain

and reset the value.  You will need to choose a shorter record

length, a greater sampling interval or fewer channels of input.

Clock Period:  Selecting this option allows the user to specify the

sampling interval at which Patchit will record data (in

milliseconds).  This value plays a role in determining the size of

the record and is limited (see Length of Record, above).  If you

choose too small a value, Patchit will complain and reset the

value.  If this happens, choose a larger value, decrease the record

length or decrease the number of input channels.

Note:  Due to the design of certain data acquisition boards, under

certain conditions Patchit may use a smaller sampling interval than

you have selected.  This should not affect the size of the memory

block described above, but the sampling interval could be smaller

than the board could handle.  This could cause the program to lock

up, a full computer reboot might then be necessary.  Always test

your parameter files before using them for critical recording

sessions!

Delay Time:  Selecting this option allows the user to specify the

delay time (in seconds) between recording sessions.  This value is

the amount of time the program waits between each record.  During

data acquisition cycles, this value is counted down on the screen.

A value of 0 will NOT cause the records to be acquired immediately

after each other.  There is a small amount of overhead time that

Patchit requires between each data acquisition cycle.

Output:  Selecting this option allows the user to control aspects

of the output channels available for use to Patchit.  There are

entries on this menu for all of the available output channels.

These can be up to 15 analog outputs (D/A converters) and one

binary output consisting of 14 separate TTL outputs.  The exact

number of outputs depends on your specific data acquisition board.

Selecting a specific output from this menu allows you to control

aspects of this output, and importantly, whether this output is

enabled for use during data acquisition.  An output will only be

used during a data acquisition cycle if it is enabled here.

Specifying outputs waveforms is described on the Sample:Stimulus

menu, below.  The values on the right side of this menu describe

some of the settings in effect for that output.  For analog outputs

("X Output X"), the enabled/disabled state of the output is shown,

and the default (holding) voltage on that output.  For binary

outputs, the enabled/disabled state of the output is shown, and the

default binary output, shown in hexadecimal.  This corresponds to

the binary state of the 14 individual TTL outputs.

X Output X:  Selecting this option allows the user to control

settings for this analog output (D/A converter) channel.

Use Channel:  Selecting this toggles the enabled/disabled state of

this output.  An output will only be used during a data acquisition

cycle if it is enabled here.

V Hold (Vh):  Selecting this allows the user to set the default

(holding) voltage on this output.  The voltage is specified in

millivolts and is emitted on this output at all times (except

during data acquisition, if the channel is enabled for output).

Note that values entered for the stimulus descriptors (see below,

Sample:Stimulus:D/A Channel X) are absolute values and are not

relative to this holding potential.

Attenuation:  Selecting this option allows the user to set the

attenuation factor for the analog output.  This value allows

Patchit to correctly set it's analog outputs to the proper values

which correspond to those needed at the external hardware (e.g. an

amplifier).  The value entered here is the factor by which

Patchit's output will be divided in order to compensate for any

attenuation or amplification at the external hardware.  For the

Dagan model 8900 amplifier which has a command potential input

which can be scaled by -10:1 or -20:1, the value entered into

Patchit would be -10 or -20 for these two scale factors,

respectively.  Generally you should use the largest attenuation

factor you can use in order to scale down any noise which might

enter your amplifier riding on the command voltage.  In my case,

using the -20:1 attenuation factor yields an effective command

voltage range at the amplifier of +/- 500 mV (for a board with D/A

converters with a range of +/- 10 V, attenuated by a factor of 20).

This value is critical and MUST be properly set in order for

Patchit to output accurate voltages.

Okay:  Selecting this causes any changes that you have made to the

state of this analog output to be saved.

Cancel:  Selecting this causes any changes that you have made to

the state of this analog output to be discarded.

Binary Output:  Selecting this option allows the user to control

the settings for the binary (TTL) output.

Use Channel:  Selecting this toggles the enabled/disabled state of

the binary output.  The binary output will only be used during a

data acquisition cycle if it is enabled here.

Default Outs:  Selecting this option allows the user to set the

default binary output state.  This is expressed as a 14 bit

hexadecimal value.  Each bit in this hexadecimal value corresponds

to the default logic state of the 14 TTL outputs.  In most cases

these outputs will be zero (0) unless you have a device whose

trigger uses negative logic (i.e. a falling edge trigger).

Okay:  Selecting this causes any changes that you have made to the

state of the binary outputs to be saved.

Cancel:  Selecting this causes any changes that you have made to

the state of this binary outputs to be discarded.

Okay:  Selecting this causes any changes that you have made to the

state of the outputs to be saved.

Cancel:  Selecting this causes any changes that you have made to

the state of the outputs to be discarded.

Inputs:  Selecting this option allows the user to control aspects

of the channels of input used during a data acquisition cycle.  The

items on this menu correspond to the available input channels (up

to 16, depending on your data acquisition board).  Selecting an

input from this menu allows you to control aspects of that input,

and importantly, whether this input is enabled.  During data

acquisition, data will only be recorded on an input if it is

enabled here.  The values on the right side of this menu describe

some of the settings in effect for this input.  The

enabled/disabled state of the input is shown, and if enabled, a

code is shown which summarizes the display settings for this input.

There are five characters between parentheses, they are, (in the

order they appear): 'D' = data will be displayed.  '1',2','3' or

'4' = The window number in which the data will be displayed.  '^'

= the data will be displayed inverted.  '*' = the displayed data

will be scaled (multiplied) by some factor.  '+' = the displayed

data will be vertically offset by some factor.  If any of these

display options are not in effect, the corresponding character will

not be shown.

Note:  Due to the design of certain data acquisition boards, under

certain conditions Patchit may have to temporarily use more input

channels than you have selected.  These unwanted channels will be

stripped away once the data has been recorded, but can have two

effects.  1) The amount of memory available for a data acquisition

cycle may be reduced.  2) The increased number of channels may

cause the sampling interval used to be smaller than you have

selected.  This sampling interval could be smaller than the board

could handle.  This could cause the program to lock up, a full

computer reboot might then be necessary.  Always test your

parameter files before using them for critical recording sessions!

X Input X:  Selecting this item allows the user to control the

settings for this input.

Recording:  Selecting this toggles the enabled/disabled state of

this output.  Data will be recorded on an input during a data

acquisition cycle only if it is enabled here.  Enabling/disabling

an input will also automatically turn on display of this data - you

can turn this off below.

Type:  This allows the user to set the type of data that is on this

channel.  This can be set to indicate that the data represents

either current or voltage.

Gain:  This allows the user to set how Patchit will handle the gain

for this data.  Four options are available.

Not Set in Patchit:  Selecting this indicates that the gain will

not be set by Patchit for this data.  If this gain option is set

for this channel and data on this channel is saved to a file and

then loaded into a data analysis program, the data analysis program

will have to prompt you for the gain.  It is usually advantageous

to set the gain in Patchit, using one of the other gain modes.

Set Gain Right Now:  Selecting this causes Patchit to prompt you

for the gain right now.  This gain mode is best suited for

circumstances in which the gain on an input is not likely to vary

often.

Prompt for Later:  Selecting this causes Patchit to prompt you for

the gain on this input channel when the data is saved to file.

This gain mode is best suited for circumstances in which the gain

might change frequently.

Read from A/D Chan:  Selecting this causes Patchit to prompt you

for the A/D channel from which automatic gain recordings will be

made.  This is presently hard-configured for my Dagan 8900

amplifier.  It has 7 gain settings -1000, -500, -200, -100, -50,

-20, -10 (with the 1G headstage) - these 7 gains generate a signal

of 0.5, 1.0, 1.5, 2.0, 2.5, 3.0 and 3.5 Volts on the gain output.

When this gain mode is active, Patchit will read the voltage from

the selected A/D gain channel prior to each data acquisition cycle.

This gain mode is best suited for amplifiers which have an

automatic gain sender unit.

Display:  Selecting this toggles the enabled/disabled state for

whether or not data on this channel will be displayed.  Data can

only be displayed if this channel is enabled for input (obviously).

Window:  This allows the user to select in which of up to four

windows data on this channel will be displayed.  If more than one

such window is available then the user will be prompted for the

window number to display the data in.  Windows are set up using

Display:Windows, above.

Note:  If data on a given channel is configured to be displayed in

a given channel, and the user then decreases the number of windows

(see Display:Windows), then any data that was to have been

displayed in a window that no longer exists (greater than the

number of windows) will be displayed in window 1.

Invert:  Selecting this toggles the enabled/disabled state for

whether or not data on this channel will be displayed inverted

(upside-down).  This feature has been implemented for amplifiers

(such as the Dagan 8900) that have a signal that is inverted (i.e.

upside-down) relative to the desired convention for that signal.

If a channel of data is configured to be displayed inverted, then

ONLY the display of this data will be inverted.  The raw data is

recorded as it is sampled from the analog signal.  If you need to

invert a channel for display, then the gain at which this data was

recorded should be a negative value.

Scale Factor:  Selecting this causes Patchit to prompt you for a

value by which to scale (multiply) the displayed data.  This

setting affects only the display of the data, the raw data is

unaffected.  Inverting the display of a channel could also be done

here by setting this value to -1, but setting invert (above) would

cause the display to be faster.

Offset:  Selecting this causes Patchit to prompt you for a value by

which to offset vertically the displayed data.  This value is in

the range +1.0 to -1.0.  +1.0 causes the data to be drawn offset

upward by half the display height, -1.0 causes the data to be drawn

offset downward by half the display height, and 0.0 causes the data

to be drawn as it is recorded.  Intermediate values cause

intermediate offsets.  This value affects only the display of the

data, the raw data is unaffected.

Okay:  Selecting this causes any changes that you have made to the

state of this input to be saved.

Cancel:  Selecting this causes any changes that you have made to

the state of this input to be discarded.

Okay:  Selecting this causes any changes that you have made to the

state of the inputs to be saved.

Cancel:  Selecting this causes any changes that you have made to

the state of the inputs to be discarded.

Stimulus:  Selecting this option allows the user to control the

output stimulus waveform(s) during data acquisition.  Patchit

controls two different types of outputs, analog (D/A) and binary

(TTL).  The stimuli that can be generated on these channels are

different and will be described separately.  Selecting one of the

options on this menu allows the user to specify the output waveform

for the selected channel.  If the output channel for which you wish

to describe the waveform is not enabled, you will be warned.

Output waveforms are described by events, these events

can occur at any time during the data acquisition cycle.  Existing

events occurring on this output channel are listed in chronological

order, in a box shown at the bottom of the screen.  Events may be

added to, changed, or deleted from this list using the menu.

Events are described by event descriptors.  The event descriptors

are different for analog and binary outputs.

X (D/A) Channel X:  This allows the user to configure the analog

output waveform for this output channel.  If this channel is not

enabled for output, you will be warned.  It can be turned on using

Sample:Outputs:X Output X:Use Channel.  The event descriptor

control menu has five items:

Add Event:  This allows the user to add an event to the event list

(and thus to the output waveform on this channel).  The event will

be added to the list in its proper chronological order.

Change Event:  This allows the user to change an existing event in

the event list (and thus the output waveform on this channel).  The

user will be prompted for which event to change.

Both Add Event and Change Event bring up the event menu:

Format:  An event can be either a step or a ramp.  A step is an

instantaneous (well, as good as the hardware can do) change in the

output state at the given time.  A ramp causes the output to change

state in a linear or ramp-like fashion starting at the given time.

The next two items on this menu differ depending upon the format

for the event that is selected:

For Step events the next two items are:

Level:  The voltage (in millivolts) to which the output is stepped

at the time of the event.

Incmt:  The voltage increment (in millivolts) by which the output

voltage is changed during each successive data acquisition cycle

(record).

For Ramp events the next two items are:

Level, Starting:  The voltage (in millivolts) at which the ramp

starts at the time of the event.

M = Slope:  The slope (in millivolts/millisecond) of the ramp.

Time:  The time (in milliseconds) at which the event is to occur.

Delta T:  The increment (in milliseconds) by which the time of the

event is shifted during each successive data acquisition cycle.

Okay:  Selecting this causes any changes that have been made to

this event to be saved.

Cancel:  Selecting this causes any changes that have been made to

this event to be discarded.

Delete Event:  This allows the user to delete an event from the

event list (and thus from the output waveform for this channel).

The user will be prompted for which event to delete.

Erase All Events:  This allows the user to delete all events in the

event list (and thus the output waveform for this channel).  If

this channel is enabled for output and there are no events for it,

then the output will be the default output voltage (set by

Sample:Outputs:X Output X:V Hold Vh).

Okay:  This exits the event descriptor control menu for this output

channel.

Binary Output:  This allows the user to configure the output

waveforms for the binary output.  If the binary output channel is

not enabled for output, you will be warned.  It can be turned on

using Sample:Outputs:Binary Output:Use Channel.  The event

descriptor control menu has five items:

Add Event:  This allows the user to add an event to the event list

(and thus to the binary output waveform).  The event will be added

to the list in its proper chronological order.

Change Event:  This allows the user to change an event in the event

list (and thus the binary output waveform).  The user will be

prompted for which event to operate upon.

Both Add Event and Change Event bring up the event menu:

Hex value:  This is the value in hexadecimal (up to 3FFF) which is

to be output on the binary port at the time of the event.  Each of

the up to 14 TTL outputs are controlled by a bit from this value.

To output the logic level "1" on binary channel 0, the value in

hexadecimal is 1.  To output the logic level "1" on binary channel

1, the value in hexadecimal is 2. Oh, go look up hexadecimal

numbers somewhere else.

Time:  The time (in milliseconds) at which the event is to occur.

Delta T:  The increment (in milliseconds) by which the time of the

event is shifted during each successive data acquisition cycle

(record).

Okay:  Selecting this causes any changes that have been made to

this event to be saved.

Cancel:  Selecting this causes any changes that have been made to

this event to be discarded.

Delete Event:  This allows the user to delete an event from the

event list and thus from the binary output waveform.  The user will

be prompted for which event to delete.

Erase All Events:  This allows the user to delete all events in the

binary output's event list (and thus the binary output waveform).

If the binary output is enabled for output and there are no events

for it, then each of the individual TTL outputs will be set to its

default (TTL) value (this is set by Sample:Outputs:Binary

Output:Default Outs).

Okay:  This exits the event descriptor control menu for the binary

output.

Okay:  This exits the Stimulus control menu.

Pulse Generator Menu:

Go:  Selecting this causes Patchit to begin acquiring data in

"pulse generator" mode.  In this mode, stimulus will be generated

on one enabled output channels and data will be acquired on one

input channel.  Simultaneously.  This process may be interrupted by

hitting the space bar or the ESC key.  The analog output for this

function is constrained to a simple square voltage pulse on the

selected output channel.  The pulse width is set at one half of the

length of the record (see below).  The voltage pulse is preceded

and anteceded by a time equal to one quarter of the length of the

record (the pulse has a 50% duty cycle).  While the pulse generator

is active, the menu bar displays the amplitude of the current pulse

and the range of pulse voltages that can be selected with the

function keys.  The initial amplitude of the pulse is +10 mV.  This

can be changed by using the function keys.  Pressing any of the

keys from F1 to F10 causes the amplitude of the pulse to be changed

on the next pulse to a voltage from +10 mV to +100 mV (F1 = 10mV,

F2 = 20mV, F3 = 30mV...).  ALT, CTRL and SHIFT modify the range of

voltages that can be selected with the function keys.  SHIFT-F1

through SHIFT-F10 generate pulses from -10 mV to -100 mV, in -10 mV

increments.  ALT-F1 through ALT-F10 generate pulses from +50 mV to

+500 mV in +50 mV increments.  CTRL-F1 through CTRL-F10 generate

pulses from -50 mV to -500 mV in -50 mV increments.  Typing the "0"

(zero) key causes the pulse to have an amplitude of 0 mV.

Length of Rec:  Selecting this option allows the user to specify

the length of the record (in milliseconds) used for the pulse

generator.  This is the duration of the each pulse generator cycle,

the pulse width is one half of this duration.  Data recorded during

the pulse generator cycle is stored in blocks known as records.

Each record is limited in size to a maximum of ~32000 points.  The

number of points is equal to (the length of the record) / (the

sampling rate).  If you choose a value for the length of a pulse

generator cycle that exceeds the maximum number of points or amount

of memory, Patchit will complain and reset this value.  You will

need to choose a shorter event length or a greater sampling

interval.

Clock Period:  Selecting this option allows the user to specify the

sampling interval (in milliseconds) at which data is recorded

during a pulse generator cycle.  This value plays a role in

determining the size of the memory block allocated and is limited

(see Length of Record, above).  If you choose too small a value,

Patchit will complain and reset the value, if this happens, choose

a larger value or decrease the length of the pulse generator cycle.

Note:  Due to the design of certain data acquisition boards, under

certain conditions Patchit may use a smaller sampling interval than

you have selected.  This should not affect the size of the memory

block described above, but the sampling interval could be smaller

than the board could handle.  This could cause the program to lock

up, a full computer reboot might then be necessary.  Always test

your parameter files before using them for critical recording

sessions!

Delay Time:  Selecting this option allows the user to specify the

delay time (in seconds) between pulse generator cycles.  This value

is the amount of time the program waits between each cycle.  A

value of 0 here will NOT cause the records to be acquired

immediately after each other.  There is a small amount of overhead

time that Patchit requires between each pulse generator cycle.

Output Channel:  Selecting this option allows the user to specify

the channel that will be used for output during a pulse generator

cycle.  It is on this channel that the voltage pulse will be

generated when Pulse Generator:Go is selected.

Input Channel:  Selecting this option allows the user to specify

the channel that will be used for input during a pulse generator

cycle.  It is on this channel that analog data from an experimental

preparation will be acquired and displayed.

Note:  Due to the design of certain data acquisition boards, under

certain conditions Patchit may have to temporarily use more input

channels than you have selected.  These unwanted channels will be

stripped away once the data has been recorded, but can have two

effects.  1) The amount of memory available for a pulse generator

cycle may be reduced.  2) The increased number of channels may

cause the sampling interval used to be smaller than you have

selected.  This sampling interval could be smaller than the board

could handle.  This could cause the program to lock up, a full

computer reboot might then be necessary.  Always test your

parameter files before using them for critical recording sessions!

Config Menu:

Hardware Config:  Selecting this option allows the user to specify

the type of data acquisition board in use and its configuration.

When this option is selected a menu appears and the user can select

among the data acquisition boards that Patchit supports.  At

present there are two families of boards that are supported, the

Data Translation DT2821 series and the Tecmar Labmaster series.

The user also has the option of selecting "No Board", this allows

you to try out Patchit on a computer that has no data acquisition

board installed.  When one of these board families is selected,

another menu appears from which the user can select the specific

type of board that is installed the hardware settings of that

board.  These options are described further below:

Family Member:  Selecting this option allows the user to specify

the exact type of data acquisition board that is installed in the

computer.  This selection allows Patchit to take advantage of the

features of that specific board (i.e. number of input channels

etc.).

Base Address:  Selecting this option allows the user to specify the

base address (in hexadecimal) of the data acquisition board that is

installed.  It is imperative that this value match that on the

board (see your board's manual).  If it is incorrectly set, Patchit

will be unable to communicate with the board.

DMA Channel(s):  Selecting this option allows the user to select

the DMA channel(s) that the program and hardware will use.  Again,

the value selected here must match the configuration of the board.

Note: At present the second DMA channel is NOT used by Patchit for

the DT2821 series boards and DMA is not used at all for Tecmar

Labmaster series boards (even the Labmaster DMA..., yes, I know, I

know...).

Interrupt:  Selecting this option allows the user to select the

interrupt channel that the program and hardware will use.  This

value should match the configuration of the board.

Note: This setting doesn't matter since Patchit doesn't use this

interrupt presently.

A/D Format:  (Labmaster Only) Selecting this option allows the user

to specify the data format used by the data acquisition board for

it's A/D converter.  This can be either Two's Complement or Binary.

This value must match the configuration of the board.  The default

format for Patchit (and Tack) is two's complement.

Trigger Src:  Selecting this option allows the user to specify

whether Patchit will use an internal or an external trigger.  For

internal trigger (the normal mode), Patchit will start its own data

acquisition cycle.  In order for this to work properly on the

DT2821 series boards the Binary Output 1, Channel 7 must be

connected to the External Trigger input.  For external trigger, a

falling-edge TTL trigger will start the data acquisition cycle.

For a DT2821 series board, the trigger should be applied to the

external trigger input.  For a Tecmar Labmaster the trigger pulse

should be applied to the binary input port B, Bit 2 (B2).

Okay:  Selecting this causes any changes that have been made to the

configuration of the data acquisition board to be saved.

Cancel:  Selecting this causes any changes that have been made to

the configuration of the data acquisition board to be discarded.

Graphics Display:  Selecting this option allows the user to specify

the type of graphics display installed and the display resolution

that Patchit should use.  Normally, the graphics autodetect setting

should be used.  For a given display type to be used, the requisite

display driver file (they have the extension ".bgi") MUST be

present in the same directory as Patchit.

Display Type:  Selecting this option allows the user to select a

graphics display type.  Twelve settings are available.

Note:  If your computer does not support the display type that you

have selected, you may well crash the program when you try to

switch display types.

Resolution:  Selecting this option allows the user to select a

display resolution for the selected display type.  Note that some

display types have only one available resolution.

Note:  If your computer's display adapter does not support the

resolution that you have selected, you may well crash the program

when you try to switch display resolutions.

Okay:  Selecting this option causes any changes that have been made

to the graphics display type and resolution to go into effect.

Cancel:  Selecting this option causes any changes that have been

made to the graphics display type and resolution to be cancelled.

Colors:  When this option is selected a color demonstration window

appears and a menu appears from which the user can select the

colors of the various components of the Patchit screen.  The

various components are the menu bar, status line, popup boxes and

data window.  For menus, the user can control the background, text,

frame and select letter colors.  For the status line, the user can

control the background and text colors.  For popup boxes, the user

can control the background, text and frame colors.  For the data

window, the user can control the background, frame and axis colors

as well as the colors for all available channels of data.

Menus:

Status Line:

Popup Boxes:

Data Window:

When the user selects one of these menu options, another menu

appears from which the user can select the color for the elements

of the windows (i.e. frame color of popup window).  Clicking on an

element will select the next available color for that element, this

will be shown in the color demonstration window.  Typing the select

letter of that element will do the same.  Holding the SHIFT key

down while typing the select letter of that element will select the

previous color for that element.  Select Okay when you are done

setting the colors.

Use Defaults:  When this option is selected, the colors of all of

the elements of the Patchit display revert to their default

settings.

Okay:  Selecting this option causes any changes that have been made

to the color scheme to go into effect.

Cancel:  Selecting this option causes any changes that have been

made to the color scheme to be cancelled.

Font:  Selecting this option allows the user to specify the font

that will be used by Patchit.  For all but the default font, the

requisite font file (they have the extension ".chr") MUST be

present in the same directory as Patchit.  An example of some text

written in the selected font is displayed.

Font:  Selecting this option allows the user to select the font

family that will be used by Patchit.  Eleven font families are

available.  The default font is a bit-mapped font, the others are

stroked (outline) fonts.

Height:  Selecting this option allows the user to set the height of

the font that will be used by Patchit.  If the selected font family

is the default, bit-mapped font, the height and the width must be

the same.

Width:  Selecting this option allows the user to set the width of

the font that will be used by Patchit.  If the selected font family

is the default, bit-mapped font, the height and the width must be

the same.

Okay:  Selecting this option causes any changes that have been made

to the font to go into effect.

Cancel:  Selecting this option causes any changes that have been

made to the font to be cancelled.

Sounds:  Selecting this option allows the user to control aspects

of the sounds that Patchit generates during its operation.  These

are described below:

Sounds:  Selecting this option toggles the state of whether Patchit

will generate sounds.  When this is disabled Patchit will make no

noise.

Type:  Selecting this option allows the user to select the type of

sound generation that Patchit will use, these are:

Single Tone:  Patchit will always generate a tone of the specified

frequency and duration (see below) for each data acquisition cycle.

Rising Scale:  When the first record is acquired, Patchit will

generate a tone of the specified frequency and duration.  For the

next and each succeeding record, the tone generated will be of a

successively higher frequency.

From A/D:  Patchit will generate a tone for each record based on

the value of a portion of the data.  When this is option is

selected the user is prompted for the channel of data from which to

record the value and is then prompted for a value from 0 to 1.

This value tells Patchit from where in the record (0 is beginning,

1 is end) to read the value of the data on which to base the sound.

Frequency:  Selecting this option allows the user to specify the

frequency (in Hz.) of the tone.  If the sound type is "Single

Tone", this is the frequency of that tone.  If the sound type is

"Rising Scale", this is the frequency generated for the first

record.  If the sound type is "From A/D", this is the frequency of

the lowest A/D value, for this sound type, the user is also

prompted for a frequency corresponding to the highest A/D value.

Duration:  When this option is selected the user is prompted for

the duration (in msec) of the tone.

Test:  When this option is selected, Patchit demonstrates the tones

that will be generated during given the currently selected options.

If the tone type is "From A/D", the demonstration plays the high

and low frequency values over and over.

Okay:  When this option is selected, you leave the sounds menu.

Display Options:  Selecting this option allows the user to control

aspects of how data is displayed on the screen.

Show all pts in record:  If this is enabled then Patchit will

display all of the data points for a given record on the screen.

That is, if there are more data points in a record than there are

pixels on the screen horizontally then Patchit will show multiple

Y points for each X point.  If this is disabled then Patchit will

only display one Y point for each X point.  Disabling this feature

causes the display to draw faster.

Connect the points:  If this is enabled then Patchit will connect

the data points on the screen with lines.  This causes the display

of data to look nicer, but significantly slows the display of data.

Auto-Erase (Storage):  If this is enabled then the screen will

automatically be cleared whenever data is acquired or whenever the

pulse generator is activated.  If this is disabled, then the screen

will not be cleared in either of these cases.

Grid and Axes:  Selecting this option allows the user to configure

how the grid and axes in the data window(s) are displayed.  There

are five items on this menu:

X Axis:  Selecting this option allows the user to specify where X

axes are to be placed in each data window.  The user can select a

variety of  X axes varying from none to three, appearing at the

top, middle and bottom of each window.

Y Axis:  Selecting this option allows the user to specify where Y

axes are to be placed in each data window.  The user can select a

variety of  Y axes varying from none to three, appearing at the

left, middle and right of each window.

Grid:  Selecting this option allows the user to specify the

appearance of the grid in the data window(s).  The grid can be

horizontal, vertical, both horizontal & vertical or neither.  The

spacing of the grid is determined by the number of divisions (see

next item).

Division:  Selecting this option allows the user to specify the

number of divisions for both the X and Y axes.  The number of

divisions determines where the ticks are placed on the X and Y axes

and the spacing of the grid.

Line Style:  Selecting this option allows the user to specify the

line style for the grid.  The line style is specified by a

hexadecimal number (0 -> FFFF) which specifies how the line is

displayed.  FFFF corresponds to a solid line, while 0 corresponds

to no line.  A value of  0F0F would make a dashed line, while 1

would make a faintly dotted line.

Okay:  Selecting this option exits the display options menu.

Other Options:  Selecting this option allows the user to control

several very important aspects of Patchit's operation.  These are

described below:

XMS Data Storage:  Enabling this feature allows Patchit to use

memory accessed via the XMS specification in order to store data

acquired before writing it to disk.  XMS memory is memory above the

1 MB boundary on PC-AT class computers (80286 and above class

microprocessors).  This memory can only be accessed if you have

installed the XMS driver HIMEM.SYS (supplied free of charge by

Microsoft), or a compatible XMS driver (e.g. Desqview's QEMM memory

manager) in your config.sys file.  The use of XMS allows Patchit to

store data above the conventional memory area, which allows more

room for Patchit to work.  This feature need not be disabled unless

you suspect a problem or else have other needs.

Startup Prm File:  Selecting this option allows the user to specify

the name of a parameter file to load whenever Patchit is started.

This feature allows the user to load a customized parameter set at

startup.  If this file is not available at startup, Patchit will

use a default set which may not be correct for your experimental

configuration! (in particular the Output Scaling will almost

certainly be wrong).

Okay:  Selecting this causes any changes that have been made to the

items on this menu to be saved.

Cancel:  Selecting this causes any changes that have been made to

the items on this menu to be discarded.

Write Config File:  Selecting this option causes the current

configuration data to be written to the Patchit.cfg configuration

file.  Only the items that are on controlled by config menu are

stored in this file.  These are:

- Hardware Settings (type of board and hardware configuration)

- Graphics display adapter & resolution

- Colors

- Font

- Sounds

- Display Settings (Show All Pts, Connect the Pts, Auto-Erase,

Grid)

- XMS Usage

- Start Up Parameter File

All other user-controllable parameters are stored in the

individual parameter (prm) files.  These values will thus be set to

the value stored in the parameter file whenever a new one is

loaded.  Of these parameters, one of the most critical (I stress

again) is the Output Scaling factor.  Please confirm that this

value is set properly in all of the parameter files which you use.

However, it may be that you would desire different output scaling

values for different parameter files, so that is why this value is

stored in the parameter file and not the configuration file.

Conclusion

I hope that you understand all of the features that I have

described here and that Patchit serves you well.  I would

appreciate your comments on the program and manual.  I always want

to know about bugs in my program and will endeavor to fix them as

speedily as is possible.  I am not responsible for any loss of data

or damage caused by this program.  I also would value your

suggestions for new features and ways to improve this program.  I

can be reached at the following addresses:

George B. Grant               (617) 625-2419 (Home Phone)

38 Banks St. #2               (617) 495-2599 (Work Phone)

Somerville, MA  02144

ggrant@husc.harvard.edu       (e-mail)

george@mander.berkeley.edu    (e-mail)

72722,3152                    (Compuserve address)

Appendix:

Hardware Configuration:

Data Translation DT2821 Family Boards:

Analog Outputs:  Patchit uses both Analog Outputs on the DT2821.

Analog Inputs:  Patchit uses all analog inputs on the DT2821.  Note

the weird channel sequence used for DT2821 family boards with less

than 16 inputs!

Binary Outputs:  Patchit uses the two Binary Output ports on the

DT2821 for binary (TTL) trigger outputs.  Two of these are put to

special use by Patchit.  Binary Output Port 1, Channel (bit) 7 and

Binary Output Port 1, Channel (bit) 6, go to logic low and logic

high, respectively, at the beginning of a record and return to

logic high and logic low, respectively, at the end of a record.

These are useful for triggering oscilloscopes.

Triggering:  If you do not need an external trigger, you should

connect the DT2821's external trigger input to Binary Output Port

1, Channel (bit) 7.  This will allow Patchit to operate properly.

If you plan to use an external trigger, it should be connected to

the external trigger input (instead of Binary Output Port 1,

Channel 7).

Tecmar Labmaster Family Boards:

General:  The following connections must be made for Patchit to

operate properly:

Source5-Out5-Start Convert

Gate5-Out4

Analog Outputs:  Patchit uses both Analog Outputs on the Tecmar

Labmaster.

Analog Inputs:  Patchit uses all analog inputs on the Tecmar

Labmaster.

Binary Outputs:  Patchit uses two of the Parallel I/O ports on the

Tecmar Labmaster for binary (TTL) trigger outputs.  These are Ports

C and A.  Patchit numbers these binary outputs from 0 to 15.

Patchit outputs 0-7 correspond to Port A, bits 0-7 and Patchit

outputs 8-15 correspond to Port C, bits 0-7.  Two of these binary

outputs are put to special use by Patchit.  Outputs 15 and 14 (C7

and C6) go to logic low and logic high, respectively,  at the

beginning of a record and return to logic high and low,

respectively, at the end of a record.  These are useful for

triggering oscilloscopes.

Note: If you have a Labmaster DMA, ports A and C come configured as

inputs (except Port C, bits 4-7).  If you want to use all sixteen

binary outputs that Patchit offers, you will need to change a

setting on the Labmaster DMA motherboard (if you choose not to do

this, you will be limited to four binary outputs, 15-12).  The

change that you need to make to the Labmaster DMA motherboard is to

reconfigure the rest of ports A and C for outputs (this change will

NOT affect pClamp).  Reseat chips USS1, USS2 and USS4 so that they

are inserted in rows 2 and 4 (the same orientation as chip USS3).

Refer to your Labmaster DMA manual for this.

Triggering:  An external trigger pulse should be applied to the

binary input port B, Bit 2 (B2).  A rising edge trigger is expected

(i.e., going from logic low to logic high).  In order to use an

external trigger, you must enable it in Patchit (see:

Config:Hardware Config:Tecmar Labmaster:Trigger Src.